## Slide 1
Problem
This project was funded by Mr. and Mrs. Charles Fox and Beyond Traditional Borders, which is made possible by a grant to Rice University from the Howard Hughes Medical Institute through the Undergraduate Science Education Program. The design work for this project was supported by the Oshman Engineering Design Kitchen.
Implementation Plans
Acknowledgements
System Criteria
Nanoparticles Harness Solar Energy
Solar-Powered Autoclave Design
Prototype Testing
Absorb entire UV, visible, and NIR spectrum
Sustain temperatures 40% higher than required for sterilization
4x more efficient than solar panels
Appropriate for resource-constrained settings
No power requirement
No noxious chemicals
Reusable
Create standalone, robust, modular sterilization system
Automate cycle
Simplify user interface
Distribute to dental and medical outreach groups that work in resource-constrained settings
Decrease risk of infection and disease transmission
Enable work in more remote areas
Conducted safety testing
Hydrostatic pressure testing
Finite element analysis
Thermal testing
Conducted full-cycle thermal testing
Achieved : 119 ̊C and 12 psig
Sterilization validated with thermally resistant bacteriological indicators
Plan to conduct additional thermal testing and durability testing.
Solar-Powered Autoclave Using Nanotechnology for the Resource-Constrained Setting
Eric Kim1, Benjamin Lu1, Kevin Schell2, Mary Quinn1, Shea Thompson1, Catherine Flaitz4, Oara Neumann3, Z. Maria Oden1, Naomi Halas3
1Dept. of Bioengineering, 2Dept. of Mechanical Engineering,3Dept. of Electrical and Computer Engineering, Rice University, Houston, TX 77005
4Dental Branch, The University of Texas Health Science Center at Houston, Houston, TX 77030
Contact: TeamNanoSPA@gmail.com
SEM Image of Gold Nanoshells
Absorption Spectrum of Gold Nanoshells
Hand
Washing
Chemical
Ethylene oxide
Heat
Autoclaving
Dry heat
Clean,
Robust
Power requirements,
Complexity
Low power,
Low temp
Toxicity,
Logistics
Simple,
Inexpensive
Does not sterilize
Advantages
Drawbacks
Current Solutions
  Our solution: Incorporate novel nanoparticle technology that efficiently absorbs solar energy to create a solar-powered autoclave
Thermal Testing Results
Fresnel lens focuses light on nanoparticle module.
1
Nanoparticle solution in module generates steam.
2
Steam moves from module into sterilizing vessel.
3
Condensed water returns to module via hydrostatic pressure.
5
Air purge system removes unsterile air from sterilizing vessel.
4
Pressure relief valve
Pressure gauge
Ball valves
Check valve
Chart: Time (min)
Goal | 115 | 115
System Criteria
Maintain 115 ̊-140 ̊C and 10-20 psi
Require only solar energy for operation
Redundant fail-safe operation
Cost < $1500
Capacity for tools required daily in mobile medical or dental clinics
Durable to transport on rugged terrain
Operation cycle < 2 hours
Achieved:
Additional testing required:
Ambient Temperature
Bottom Thermocouple
Goal Temperature
Rural dental and medical clinics require means of sterilizing tools for procedures.
No sterilization technique is optimized for resource-constrained settings.
  Limited power supply and funds
  Unreliable supply chain
  Untrained technicians
Improper sterilization results in increased risk of infection within developing countries.
  Goal: To develop a cost effective, robust, and portable solar-powered autoclave for resource-constrained settings