## Slide 1
OxCart: Oxygen Concentrator and Analyzer for the Developing World
Matt Amdahl, Ryan Le, Dan Nelson, Jay Patel, Abe Segura
Department of Bioengineering, Rice University, Houston TX 77005
Team RedOx <redoxconcentrators@gmail.com>
Objective
Concentrator Design Objectives
To design and build:
- A low-purity, high-flow oxygen concentrator. - A reliable, inexpensive oxygen analyzer.
Criteria | Goals
Oxygen Flow | > 15 L/min, 40% O2
Power Consumption | < 400W
Longevity | > 5 years
Size | < 2.5’ x 2.5’ x 8’
Cost | < $1000
Motivation for OxCart
Sensor Design Objectives
Criteria | Goals
Range | 21 – 70% O2
Accuracy | < 5% absolute error
Power Demand | < 10 W
Cost | < $50
Concentrator Design, Testing and Results
Analyzer Design, Testing and Results
Conclusions
Air
O2
Air
Closed
Air
Air
Pressurization Time (seconds) | Purge Time (seconds) | Maximum %O2 Reached (%) | Average Pressure (psi)
30 | 10 | 40.1 | 16.3
30 | 15 | 38.0 | 18.3
45 | 15 | 41.0 | 26.0
40 | 10 | 34.9 | 23.1
5 | 5 | 29.9 | 5.0
Closed
Acknowledgements & References
We would like to thank Dr. Maria Oden, Dr. David Hilmers, Dr. John Graf, Dr. Gary Woods, Carlos Amaro, Joe Gesenhus the Oshman Engineering Design Kitchen (OEDK), Rice University’s Dept. of Bioengineering, Beyond Traditional Borders (BTB), NASA, Wyle Labs, Lockheed Martin, and the Bioastronautics Contract.
Duke, T., et al. Improved oxygen systems for childhood pneumonia: a multihospital effectiveness study in Papua New Guinea. Lancet, 2008; 372:1328-33
Howie, S.R.C., et al. Meeting oxygen needs in Africa: an options analysis from the Gambia. Bull World Health Organ, 2009; 87:763-771.
Achievements
  Prototype testing indicates both systems are
  capable of achieving the design objectives.
Recommendations
  A top-top equalization step might increase the
  efficiency of the concentrator, helping decrease cost.
  The analyzer may be modified to rely on smaller,
  less expensive button zinc-air batteries.
Benefits
  The OxCart will offer healthcare providers a cheaper
  and more reliable respiratory treatment option than
  current standards.
2 million children die each year from acute respiratory illness.
  98% of all deaths occur in developing regions.1
Oxygen therapy reduces mortality from pneumonia by 35%.2
Current oxygen delivery methods in developing areas are
  Prohibitively expensive (upwards of $1000)
  Prone to malfunction
  Fail to measure oxygen concentration
Current portable oxygen analyzers cost upwards of $200.
Stand-alone component
Relies on existing zinc-air battery technology
Battery voltage dependent on ambient oxygen concentration
Step | Left Bed | Right Bed
1 | Venting/Purging | Pressurizing
2 | Venting/Purging | Producing
3 | Pressurizing | Venting/Purging
4 | Producing | Venting/Purging
(Top) Fig 3. CAD drawing of OxCart.
(Right) Fig 4. Current prototype of concentrator
The OxCart uses a Pressure Swing Adsorption (PSA) system:
System contains several beds full of 5A Zeolite
5A Zeolite preferentially adsorbs nitrogen when pressurized
  Thus purifying the air, allowing for higher
  concentrations of oxygen
Depressurizing and purging zeolite regenerates it.
Varying cycle time leads to varying %O2 concentrations.
As cycle time increases past a certain point, the output oxygen purity begins to decrease, but oxygen recovery will increase.
We are testing to find the maximum oxygen purity attained at each cycle time, time it takes to reach this maximum purity, and output flow
O2
Fig 2. The Four steps of a PSA cycle.
Table 1. Bed states during each step of PSA cycle.
(Top) Graph 2. Comparison of our analyzer to commercial sensor
(Right) Fig 1. Circuit diagram of analyzer
Table 2. The data represented here is only sample data. Multiple tests under similar pressurization and purge times have been run. However, there are varying results because of varying qualities of zeolite as well as pressure drops.
3 ways to relate battery voltage to O2 concentration
  Difference between final and baseline signal
  Difference between final and initial signal
  Amplified differential signal between 2 batteries
  Analyzer calibrated and tested between 21% - 93% O2
Logarithmic model relates voltage and O2 concentration
Initial tests show analyzer capable of accuracy within 1%
Graph 1. The graph shows how closely our data for the differential voltage signal between two batteries correlates with a best fit log line.
Chart: Calculated O2 Concentration vs. Actual O2 Concentration
| 21.0 | 25.0 | 30.0 | 35.0 | 40.0 | 45.0 | 50.0 | 55.0 | 60.0 | 65.0 | 70.0
 | 20.5171 | 25.1859 | 29.256 | 36.0821 | 41.3088 | 45.3112 | 50.1689 | 53.923 | 58.0796 | 63.1485 | 65.4747
Chart: Differential Voltage vs. %O2
| 70.0 | 65.0 | 60.0 | 55.0 | 50.0 | 45.0 | 40.0 | 35.0 | 30.0 | 25.0 | 21.0
 | 5.64 | 5.3033 | 5.03 | 4.7567 | 4.31 | 3.8367 | 3.3267 | 2.8133 | 2.1467 | 1.33 | 0.56333